BOREAS HYD-09 Belfort Rain Gauge Data

Summary

This data set contains the measurements from the Belfort rain gauges at the BOREAS NSA andSSA.  These measurements were submitted in 15-minute and 1-hour intervals.  Only the 15-minute interval data set was loaded into the data base tables.  Data were collected from the Belfort gauges from mid-April until mid-October in 1994, 1995, and 1996.  The data are available as tabular ASCII files.

Table of Contents

*  1 Data Set Overview

*  2 Investigator(s)

*  3 Theory of Measurements

*  4 Equipment

*  5 Data Acquisition Methods

*  6 Observations

*  7 Data Description

*  8 Data Organization

*  9 Data Manipulations

*  10 Errors

*  11 Notes

*  12 Application of the Data Set

*  13 Future Modifications and Plans

*  14 Software

*  15 Data Access

*  16 Output Products and Availability

*  17 References

*  18 Glossary of Terms

*  19 List of Acronyms

*  20 Document Information

1. Data Set Overview

1.1 Data Set Identification

BOREAS HYD-09 Belfort Rain Gauge Data

1.2 Data Set Introduction

This data set contains the measurements from the Belfort rain gauges at the BOReal Ecosystem-Atmosphere Study (BOREAS) Northern Study Area (NSA) and Southern Study Area (SSA).  These measurements were submitted in 15-minute and 1-hour intervals.  Only the 15-minute interval data set was loaded into the data base tables.

1.3 Objective/Purpose

This investigation will seek to identify, through field measurements and computer modeling, the space-time distribution of meltwater supply to the soil during the spring melt period and the evolution of soil moisture, evaporation, and runoff from the end of the snowmelt period through freeze up.  The snow modeling activity will consist of two components. The first phase will use existing "off- the-shelf" models to forecast the onset and spatial extent of snowmelt and meltwater supply to the soil column prior to the 1994 Intensive Field Campaigns (IFCs).  The second phase will extend, implement, and verify a physically based energy balance snowmelt model of the two sites and will evaluate approaches to aggregating detailed snowmelt predictions and measurements based on the model to larger scales, up to the size of a rectangle of several hundred km containing the northern and southern sites. The soil moisture modeling is based on a grouped response unit method tat will allow characterization of soil moisture, evaporation, and runoff for the entire northern and southern sites.

1.4 Summary of Parameters

The following phenomena and their parameters are being reported:

Precipitation amount in a 15-minute period.

1.5 Discussion

The locations for 12 tipping bucket measuring devices, 10 Belfort gauges, and 5 stream sites were selected within the two BOREAS study sites (NSA and SSA). These instruments were installed during the 1994 Focused Field Campaign-Thaw (FFC-T) in the last week of April.  They were in operation until October 1994 when they were removed from service. The tipping buckets and Belfort gauges provided an approximate measure of the precipitation in the study areas and the discharge rates of the streams provided a measurement of water leaving the study area.  When used together, these two sets of data provide a balance of the water cycle.

Similar measurements were collected in 1995 and 1996 from approximately April to October of those years.

1.6 Related Data Sets

BOREAS HYD-09 Tipping Bucket Rain Gauge Data

BOREAS HYD-09 Stream Flow Data

2. Investigator(s)

2.1 Investigator(s) Name and Title

Prof. Ric Soulis

University of Waterloo

Department of Civil Engineering

2.2 Title of Investigation

From Micro-Scale to Meso-Scale Snowmelt, Soil Moisture and Evapotranspiration from Distributed Hydrologic Models

2.3 Contact Information

Contact 1

-----------------------

Dr. Nicholas Kouwen

Univ. of Waterloo

Waterloo, ON

(519) 885-1211 x3309

(519) 888-6197

kouwen@sunburn.uwaterloo.ca

Contact 2

-----------------------

Dr. Ric Soulis

Univ. of Waterloo

Waterloo, ON

(519) 885-1211 x2175

(519) 888-6197

ric@sunburn.uwaterloo.ca

Contact 3

------------------------

David Knapp

NASA/GSFC

Greenbelt, MD

(301) 286-1424

David.Knapp@gsfc.nasa.gov

3. Theory of Measurements

The Belfort gauges, used to measure precipitation, were 4 ft tall, had an orifice of 8 inches in diameter on top, and were supported in a stable, upright position on a wooden base at all of the sites.  They are designed to find the vertical depth of water that would accumulate on a level surface if the precipitation remained where it fell.

Precipitation falls through the 8-inch opening at the top of the apparatus and is funneled into a bucket near its center. This bucket is mounted on a gravimetric measuring device. Different weights are transmitted to the data logger as different voltages and recorded for later use.  When the data are retrieved, the voltages must be converted to the corresponding volumes and precipitation using calibration coefficients.

4. Equipment

4.1 Sensor/Instrument Description

Belfort Gauge:

A Belfort gauge is a device used to measure the amount of precipitation in a given area.  When rain falls, it is collected in a bucket and weighed.   When the weight of the water is calculated and compared to the calibration coefficient, the amount of rainfall can be calculated.  The data logger stores this information at regular intervals until it can be retrieved.

The Belfort gauges used were originally equipped with chart recorders.  They were converted to electronic recorders before they were installed using the linear potentiometer so that data could be recorded at 15-minute intervals.

Linear Potentiometer:

The LP-XXFP linear potentiometer is a device designed to measure differences in resistance.  The resistance is varied by a movable shaft.  When the shaft is moved by the different weights of the collecting bucket, the resistance is changed, resulting in different voltages.  The volume of water can then be calculated.

Data Logger:

The chart pac CP-X data logger was connected to the Belfort gauge so that it could record water levels every 15 minutes. The data logger stored the information in its memory until it was retrieved.

4.1.1 Collection Environment

The equipment was operated in variable ambient atmospheric conditions at the field sites in Saskatchewan and Manitoba during 1994.

4.1.2 Source/Platform

The Belfort gauge was mounted on a wide wooden base placed on the ground.

The linear potentiometer was mounted inside the Belfort gauge housing.

The data logger was mounted inside the Belfort gauge housing.

4.1.3 Source/Platform Mission Objectives

In order to make a detailed estimate of the total amount of precipitation, sufficient precipitation data must be collected to account for all peak periods of precipitation and any variations that occur over time. To achieve this, the data were recorded at regular intervals of 15 minutes.  The data logger stored the information until it was retrieved.  Tipping bucket and Belfort gauges were used to provide independent measures of precipitation so that the error would be reduced.

4.1.4 Key Variables

Belfort Gauge - Precipitation amount in a 15-minute period

Linear Potentiometer - Voltage

Data Logger - Voltage

4.1.5 Principles of Operation

The Belfort gauge is a self-contained device designed to measure the amount of precipitation occurring in a given area over time.  It is only an estimate of the total rainfall in an area.  The amount of water is measured by weighing a bucket filled with rainwater and converting the weight to millimeters of rainfall using a calibration coefficient.

4.1.6 Sensor/Instrument Measurement Geometry

The Belfort gauge is mounted at ground level.  It must be in a clearing large enough so that a 45-degree inverted cone originating from its highest point can reach the atmosphere unimpeded.  If trees or other obstructions are surrounding the Belfort gauge, they will affect the data.  The device should be mounted on a stable base (example: a wooden platform) and leveled so that the opening is completely vertical (i.e., axis of gauge is perpendicular to a water level).

4.1.7 Manufacturer of Sensor/Instrument

Belfort Gauge

Belfort Instrument Company

727 S. Wolfe Street

Baltimore, MD,  21231  U.S.A.

Linear Potentiometer

MIDORI America Corporation

Corona CA, U.S.A.

Data Logger - Lakewood Systems

Edmonton, Alberta, Canada

(403) 462-9211

4.2 Calibration

A Belfort gauge is calibrated by pouring a known amount of water into the orifice at the top of the gauge or adding the calibration weights equivalent to known quantities of water and comparing this with the weight that the gauge registers.  This will allow for the calculation of the millimeters of precipitation per milligram of weight.  It is possible to find the volume of water associated with a certain mass by using the specific density of water at that temperature.  The volume of water can then be compared to the specification sheet and a calibration coefficient calculated.

4.2.1 Specifications

Belfort Gauge - orifice opening: 8 inches

bucket size: 12 quarts

orifice height above ground: 4 ft

Linear Potentiometer - linearity: +/- 1%

friction: 40-g maximum

operating temperatures: -25°C to 80°C

4.2.1.1 Tolerance

A Belfort gauge can measure volumes of water from 0 to 12 quarts (the volume of the bucket).  The gauge can measure precipitation in 0.01-inch increments.

4.2.2 Frequency of Calibration

The Belfort gauges were calibrated once at the beginning, the middle, and the end of the study period.

4.2.3 Other Calibration Information

Not applicable.

5. Data Acquisition Methods

A data logger recorded the precipitation data for the Belfort gauges.  The data logger was connected to a notebook computer onsite, where the logger's stored information was transferred to the notebook.

6. Observations

6.1 Data Notes

The following table describes the problems with the Belfort gauges.

| Year | Date | Event |
| --- | --- | --- |
| Belfort 1 |
| 1994 | 10-Apr | Bel1 installed at Hwy 913 & 120 53° 51' 39"N 105° 56' 27"W approx. |
|  | 05-May | Bel1 moved to White Gull Drain |
|  | 11-May | Gauge reset at 11:00 |
|  | 21-Jul | Gauge knocked over |
|  | 24-Jul | Lid knocked off |
|  | 16-Oct | Bel1 removed |
| 1995 | 26-Apr | Bel1 installed |
|  | 17-Jun - 06-Jul | Missing data - mechanical problems |
|  | 07-Nov | Bel1 removed |
| 1996 | 05-Jun | Bel1 installed |
|  | 06-Nov | Bel1 removed |
| Belfort 2 |
| 1994 | 10-Apr | Bel2 installed |
|  | 01-Jun | Gauge knocked over |
|  | 13-Jun | Lid knocked off, affecting all subsequent readings; readings adjusted by using the ratio of the diameter of the gauge opening to the diameter of the bucket |
|  | 17-Oct | Bel2 removed |
| 1995 | 12-Apr | Bel2 installed |
|  | 12-Apr - 26-Apr | Precip recordings by TB5 |
|  | 30-Aug | Water level lowered - Belfort gauge hit maximum water level |
|  | 08-Nov | Bel2 removed |
| 1996 | 15-Apr | Bel2 installed |
|  | 06-Nov | Bel2 removed |
| Belfort 3 |
| 1994 | 11-Apr | Bel3 installed |
|  | 28-Jun | Lid knocked off gauge |
|  | 21-Jul - 28-Jul | Missing data |
|  | 14-Oct | Gauge damaged by bear |
|  | 17-Oct | Bel3 removed |
| 1995 | 27-Apr | Bel3 installed |
|  | 07-Nov | Bel3 removed |
| 1996 | 05-Jun | Bel3 installed |
|  | 24-Jul - 11-Sep | Missing data |
|  | 05-Nov | Bel3 removed |
| Belfort 4 - Tipping Bucket 9 Note: Was replaced with TB9 for 1995 and 1996 |
| 1994 | 21-Apr | Bel4 installed |
|  | 13-Jun | Lid knocked off |
|  | 17-Oct | Bel4 removed |
| 1995 |  | Installed as TB9 - see Tipping Bucket information |
| 1996 |  | Installed as TB9 - see Tipping Bucket information |
| Belfort 5 |
| 1994 | 12-Apr | Bel5 installed |
|  | 04-May - 09-May | Gauge knocked over during this period, data unreliable |
|  | 16-May | Gauge knocked over during windstorm |
|  | 17-May | Gauge reset |
|  | 16-Oct | Bel5 removed |
| 1995 | 11-Apr | Bel5 installed |
|  | 07-Nov | Bel5 removed |
| 1996 | 23-Aug | Bel5 installed |
|  | 06-Nov | Bel5 removed |
| Belfort 21 - Tipping Bucket 26 Note: Was replaced with TB26 for 1995 and 1996 |
| 1994 | 11-May | Bel21 installed |
|  | 12-May | Logger changed to double precision |
|  | 11-Jul - 12-Jul | Gauge knocked over, repaired and reset; missing data |
|  | 04-Oct | Bel21 removed |
| 1995 |  | Installed as TB26 - see Tipping Bucket information |
| 1996 |  | Installed as TB26 - see Tipping Bucket information |
| Belfort 22 - Tipping Bucket 28 Note: Was replaced with TB28 for 1995 and 1996 |
| 1994 | 29-Apr | Bel22 installed |
|  | 09-Jul | Changed logger number to 30 from 28 |
|  | 13-Oct | Bel22 removed |
| 1995 |  | Installed as TB28 - see Tipping Bucket information |
| 1996 |  | Installed as TB28 - see Tipping Bucket information |
| Belfort 23 - Tipping Bucket 27 Note: Was replaced with TB27 for 1995 and 1996 |
| 1994 | 28-Apr | Bel23 installed |
|  | 30-Jul | Gauge knocked over |
|  | 02-Aug | Gauge repaired |
|  | 13-Oct | Bel23 removed |
| 1995 |  | Installed as TB27 - see Tipping Bucket information |
| 1996 |  | Installed as TB27 - see Tipping Bucket information |
| Tipping Bucket 21 - Belfort 24 Note: Was TB21 in 1994 replaced with Bel24 in 1995 and 1996 |
| 1994 |  | Installed as TB21 - see Tipping Bucket information |
| 1995 | 23-Apr | Bel24 installed |
|  | 10-Nov | Bel24 removed |
| 1996 | 13-Apr | Bel24 installed |
|  | 04-Jul - 02-aug | Missing data |
|  | 06-Nov | Bel24 removed |
| Tipping Bucket 24 - Belfort 25 Note: Was TB24 in 1994 replaced with Bel25 in 1995 and 1996 |
| 1994 |  | Installed as TB24 - see Tipping Bucket information |
| 1995 | 23-Apr | Bel25 installed |
|  | 10-Nov | Bel25 removed |
| 1996 | 13-Apr | Bel25 installed |
|  | 06-Nov | Bel25 removed |
| Belfort at NW2 |
| 1994 | 25-Apr | Bel NW2 installed |
|  | 14-Oct | Bel NW2 removed |
| 1995 | 22-Apr | Bel NW2 installed |
|  | 03-Nov | Missing data |
|  | 09-Nov | Bel NW2 removed |
| 1996 | 13-Apr | Bel NW2 installed |
|  | 23-Oct | Bel NW2 removed |
| Belfort at NW3 |
| 1994 | 25-Apr | Bel NW3 installed |
|  | 14-Oct | Bel NW3 removed |
| 1995 | 23-Apr | Bel NW3 installed |
|  | 09-Nov | Bel NW3 removed |
| 1996 | 13-Apr | Bel NW3 installed |
|  | 09-Jun | Bel NW3 moved downstream and renamed NW3A |
|  | 24-Oct | Bel NW3A removed |

6.2 Field Notes

None given.

7. Data Description

7.1 Spatial Characteristics

Ten Belfort gauges were set up in and around the boundaries of the water basins that were being studied.  They gave an estimation of the precipitation within these boundaries.  Each Belfort gauge was a point measurement of the precipitation.  From these point measurements, the average precipitation for the study basins could be estimated.

7.1.1 Spatial Coverage

Each Belfort gauge is a point measurement of the precipitation, making their total area of coverage very small.  Together they represent the total rainfall in the area.  The size of both northern water basins being studied is 27 km2, while the size of the southern water basin is 574 km2.

The following table lists the Belfort gauges used and their locations.

These coordinates are in the North American Datum 1983 (NAD83).  In some cases, gauges were moved during the measurement campaigns.  If the change in location was significant, the new location was given a new site identifier.  Information regarding changes in location and servicing of the sites is listed in Section 11.2.

SITE_ID                LONGITUDE    LATITUDE

--------------------  ----------  ----------

NSA-B21-HYD09-BLRG21   98.41114W   55.88857N

NSA-B22-HYD09-BLRG22   98.49892W   55.77653N

NSA-B23-HYD09-BLRG23   98.55822W   55.89147N

NSA-B24-HYD09-BLRG24   98.34275W   55.91236N

NSA-B25-HYD09-BLRG25   98.566W     55.92661N

NSA-BN2-HYD09-BLRGN2   98.52806W   55.91528N

NSA-BN3-HYD09-BLRGN3   98.37603W   55.91686N

SSA-BL1-HYD09-BLRG01  105.01149W   53.89156N (moved slightly 05-May-1994)

SSA-BL2-HYD09-BLRG02  105.03775W   54.08506N

SSA-BL3-HYD09-BLRG03  104.76263W   54.08551N

SSA-BL4-HYD09-BLRG04  104.89617W   54.0047N

SSA-BL5-HYD09-BLRG05  104.82002W   53.92625N

7.1.2 Spatial Coverage Map

Not available.

7.1.3 Spatial Resolution

The resolution of the resultant data is low because the data were gathered as point measurements.  They are meant only to give the average precipitation rates at 15-minute intervals for the study basins.

7.1.4 Projection

Not applicable.

7.1.5 Grid Description

Not applicable.

7.2 Temporal Characteristics

7.2.1 Temporal Coverage

Data were collected from the Belfort gauges from mid-April until mid-October in 1994, 1995, and 1996.

7.2.2 Temporal Coverage Map

Not available.

7.2.3 Temporal Resolution

Measurements were recorded by the Belfort gauges at 15-minute intervals.

7.3 Data Characteristics

Data characteristics are defined in the companion data definition file (h9rgbl94.def).

7.4 Sample Data Record

Sample data format shown in the companion data definition file (h9rgbl94.def).

8. Data Organization

8.1 Data Granularity

All of the Belfort Rain Gauge Data are available by the year.

8.2 Data Format(s)

The data files contain numerical and character fields of varying length separated by commas. The character fields are enclosed with  a single apostrophe marks. There are no spaces between the fields.  Sample data records are shown in the companion data definition files (h9rgbl94.def).

9. Data Manipulations

9.1 Formulae

After a Belfort gauge was calibrated, the amount of precipitation per measuring period was found by using:

Total precipitation = weight of accumulated water * equivalent precipitation per milligram of weight

i.e.,  4.5 mm = 45 milligrams * 0.1 mm of rain per milligram

9.1.1 Derivation Techniques and Algorithms

At regular intervals during the monitoring period, the accumulated weight of water in the bucket was recorded on a data logger.  The weight was converted first to a volume of water and then to precipitation using the calibration coefficient.  If the Belfort gauge was reset (bucket was emptied and data logger started to record from zero again), then the data were adjusted accordingly.

Temperature fluctuations caused some variation in the collected data.  An algorithm was run on the data to correct for this error.

9.2 Data Processing Sequence

9.2.1 Processing Steps

The following steps were performed to collect the data:

1)	Set up necessary equipment.

2)	Measure the accumulated weight of water over time.

3)	Calculate the cumulative precipitation amounts using the coefficient.

4)	Convert cumulative amounts to incremental amounts and write ASCII files with the appropriate identifying information noted beside each row (location, year, day, month).

5)	Add the necessary column headings.

6)	Transfer the information to the data base.

BOREAS Information System (BORIS) Staff processed the data by:

1)	Reviewing the initial data files and loading them online for BOREAS team access.

2)	Designing relational data base tables to inventory and store the data.

3)	Loading the data into the relational data base tables.

4)	Working with the Hydrology (HYD-09) team to document the data set.

5)	Extracting the standardized data into logical files.

9.2.2 Processing Changes

None.

9.3 Calculations

9.3.1 Special Corrections/Adjustments

None.

9.3.2 Calculated Variables

None.

9.4 Graphs and Plots

See associated report file HYD9REP.DOC.

10. Errors

10.1 Sources of Error

Most of the errors occur during the actual measuring of the precipitation.  Some of the smaller sources of error occur from water splashing out of the funnel, evaporating water that cannot be measured,  water used to initially wet the gauge's funnel and inside surfaces, and gauges that are not perfectly level.

A larger source of error is the wind which can  cause turbulent air flow around the gauge, creating updrafts and downdrafts that interfere with the normal path of precipitation.  A high wind speed will create compressed lines of flow, reducing the amount of precipitation that enters the gauge.  The higher the wind speed, the greater the effect on the measured precipitation.

If the collecting rim (opening to the outside) is damaged in any way, the amount of precipitation being measured will be changed because less water can enter the gauge than it was calibrated for.

10.2 Quality Assessment

10.2.1 Data Validation by Source

The effects of the wind were reduced considerably by placing the gauges in clearings surrounded by forest which shielded the gauge from oncoming winds.  Evaporation was reduced through the use of a thin film of oil placed in the bucket.

After the information was collected, it was run through an algorithm to detect any data that were abnormal when compared to the rest of the data.  Computer programs were also used to adjust the gauges for periods when they had missing lids.

There may be anomalies in the data that could not be resolved at the time this data was published.  Please consult the data notes (Section 6.1) for information that might give some indication about the source of any anomalies.

10.2.2 Confidence Level/Accuracy Judgment

The confidence level of the data varies with the particular Belfort gauge at the time of measurement.

10.2.3 Measurement Error for Parameters

The collected data will eventually be compared to radar data and the calibration measurements, but at this time no steps have been taken to carry this out.

10.2.4 Additional Quality Assessments

None.

10.2.5 Data Verification by Data Center

Data that were loaded into the data tables were spot checked against the original ASCII data that were submitted to check for data loading errors.

11. Notes

11.1 Limitations of the Data

None given.

11.2 Known Problems with the Data

See Section 6.1

11.3 Usage Guidance

Because of the problems that occurred, some periods of precipitation are not accurate.

During the study period, the temperature was changing which affected the observations of the measuring equipment.  Although this was corrected with an algorithm, the data were not as accurate those obtained from a float tube at a constant temperature.

11.4 Other Relevant Information

HYD-09 wrote a report that is available through its Web page.  See Section 9.4 for details.

Kouwen, N., R. Soulis, W. Jenkinson, A. Graham and T. Neff.  1997. BOREAS: Boreal Forest Hydrological Research Study.  Hydrology 9 Group: From Micro-scale to meso-scale snowmelt, soil moisture and evapotranspiration from distributed hydrological models,  University of Waterloo, Dept. of Civil Engineering, August 1997.

12. Application of the Data Set

None given.

13. Future Modifications and Plans

None.

14. Software

Several computer programs are required to convert the data logger information into precipitation and to check it for quality.

14.1 Software Description

None given.

14.2 Software Access

None given.

15. Data Access

15.1 Contact Information

Ms. Beth Nelson

BOREASData Manager

NASA/GSFC

Greenbelt, Maryland

(301) 286-4005

(301) 286-0239 (fax)

beth@ltpmail.gsfc.nasa.gov

15.2 Data Center Identification

See Section 15.1.

15.3 Procedures for Obtaining Data

Users may place requests by telephone, electronic mail, or fax.

15.4 Data Center Status/Plans

The HYD-09 Belfort rain gauge data are available from the Earth Observing System Data and Information System (EODIS) Oak Ridge National Laboratory (ORNL) Distributed Active Archive Center (DAAC).

The BOREAS contact at ORNL is:

ORNL DAAC User Services

Oak Ridge National Laboratory

Oak Ridge, TN

(423) 241-3952

ornldaac@ornl.gov

ornl@eos.nasa.gov

16. Output Products and Availability

16.1 Tape Products

None.

16.2 Film Products

None.

16.3 Other Products

The data and documentation are available as ASCII files.

17. References

17.1 Platform/Sensor/Instrument/Data Processing Documentation

Belfort Instrument Company.  Instruction Manual Catalog Number 5-780 Series Universal Recording Rain Gauge.  1986.

Hoskin Scientific Limited.  1992.  Chart Pac Cp-X (price and specification sheet).

Kouwen, N., R. Soulis, W. Jenkinson, A. Graham and T. Neff. 1997. BOREAS: Boreal Forest Hydrological Research Study.  Hydrology 9 Group: From Micro-scale to meso-scale snowmelt, soil moisture and evapotranspiration from distributed hydrological models,  University of Waterloo, Dept. of Civil Engineering, August 1997.

MIDORI America Corporation, Corona, CA.  Specification Diagram for Model LP-XXFP Linear Potentiometer.

17.2 Journal Articles and Study Reports

Sellers, P. and F. Hall. 1994. Boreal Ecosystem-Atmosphere Study: Experiment Plan. Version 1994-3.0, NASA BOREAS Report (EXPLAN 94).

Sellers, P. and F. Hall. 1996. Boreal Ecosystem-Atmosphere Study: Experiment Plan. Version 1996-2.0, NASA BOREAS Report (EXPLAN 96).

Sellers, P. and F. Hall. 1997. BOREAS Overview Paper. JGR Special Issue (in press).

Sellers, P., F. Hall, and K.F. Huemmrich. 1996. Boreal Ecosystem-Atmosphere Study: 1994 Operations. NASA BOREAS Report (OPS DOC 94).

Sellers, P., F. Hall, and K.F. Huemmrich. 1997. Boreal Ecosystem-Atmosphere Study: 1996 Operations. NASA BOREAS Report (OPS DOC 96).

Sellers, P., F. Hall, H. Margolis, B. Kelly, D. Baldocchi, G. den Hartog, J. Cihlar, M.G. Ryan, B. Goodison, P. Crill, K.J. Ranson, D. Lettenmaier, and D.E. Wickland. 1995. The boreal ecosystem-atmosphere study (BOREAS): an overview and early results from the 1994 field year. Bulletin of the American Meteorological Society. 76(9):1549-1577.

17.3 Archive/DBMS Usage Documentation

None.

18. Glossary of Terms

None.

19. List of Acronyms

AES     - Atmospheric Environment Service of Canada

ASCII   - American Standard Code for Information Interchange

BEL     - Belfort Gauge

BOREAS  - BOReal Ecosystem-Atmosphere Study

BORIS   - BOREAS Information System

CD-ROM  - Compact Disk (optical), Read-Only Memory

DAAC    - Distributed Active Archive Center

EOS     - Earth Observing System

EOSDIS  - EOS Data and Information System

FFC-T   - Focused Field Campaign - Thaw

GMT     - Greenwich Mean Time

GSFC    - Goddard Space Flight Center

HYD     - Hydrology

IFC     - Intense Field Campaign

ml     - milliliters

mm     - millimeters

NAD83   - North American Datum 1983

NASA    - National Aeronautics and Space Administration

NSA     - Northern Study Area

ORNL    - Oak Ridge National Laboratory

PANP    - Prince Albert National Park

SSA     - Southern Study Area

TB      - Tipping Bucket

URL     - Uniform Resource Locator

20. Document Information

20.1 Document Revision Dates

Written:         02-Dec-1996

Last Updated:    26-Feb-1998

20.2 Document Review Dates

BORIS Review:   12-Sep-1997

Science Review: 01-Oct-1997

20.3 Document ID

20.4 Citation

The efforts of Ric Soulis and Nick Kouwen from the University of Waterloo in collecting these data and making them available as well as the efforts of BORIS staff in processing the data are greatly appreciated.

20.5 Document Curator

20.6 Document URL

Keywords

-------------

RAINFALL