BOREAS RSS-12 Automated Ground Sunphotometer Measurements in the SSA

Summary

The BOREAS RSS-12 team collected both ground and airborne sunphotometer measurements for use in characterizing the aerosol optical properties of the atmosphere during the BOREAS data collection activities.  These measurements are to be used to: 1) measure the magnitude and variability of the aerosol optical depth in both time and space; 2) determine the optical properties of the boreal aerosols; and 3) atmospherically correct some remotely sensed data acquired during BOREAS. These data cover selected days and times from May to September 1994 and were taken from one of two ground sites near Candle Lake in the SSA.  The data described in this document are from the field sunphotometer data. The data are stored in tabular ASCII files.

Table of Contents

*  1 Data Set Overview

*  2 Investigator(s)

*  3 Theory of Measurements

*  4 Equipment

*  5 Data Acquisition Methods

*  6 Observations

*  7 Data Description

*  8 Data Organization

*  9 Data Manipulations

*  10 Errors

*  11 Notes

*  12 Application of the Data Set

*  13 Future Modifications and Plans

*  14 Software

*  15 Data Access

*  16 Output Products and Availability

*  17 References

*  18 Glossary of Terms

*  19 List of Acronyms

*  20 Document Information

1. Data Set Overview

1.1 Data Set Identification

BOREAS RSS-12 Automated Ground Sunphotometer Measurements in the SSA

1.2 Data Set Introduction

The Automated Ground Sunphotometer (AGSP) data set consists of instrument voltages; Sun position information; and ozone (O3), nitrogen dioxide (NO2) and aerosol optical depth values.  These data were collected and processed by the BOReal Ecosystem-Atmosphere Study (BOREAS) Remote Sensing Science Team 12 (RSS-12) at the National Aeronautics and Space Administration (NASA) Ames Research Center (ARC).  The data provide a good characterization of atmospheric aerosols during the data collection periods.

1.3 Objective/Purpose

The overall goal of this investigation was to measure aerosol optical properties from both ground- and aircraft-based sunphotometers during the BOREAS Intensive Field Campaigns (IFCs). These measurements are to be used to:

1) Measure the magnitude and variability of the aerosol optical

depth in both time and space.

2) Determine the optical properties of the boreal aerosols.

3) Atmospherically correct selected remotely sensed data acquired

during BOREAS.

1.4 Summary of Parameters

The phenomenon being measured is the atmospheric aerosol optical depth. The parameters include Rayleigh optical depth, aerosol optical depth, time, latitude, longitude, air mass, and solar position.

1.5 Discussion

The AGSP data will be used in conjunction with the Airborne Tracking Sunphotometer (ATSP) data to determine the magnitude and variability of the aerosol optical depth in both time and space.  The aerosol optical depth data will be inverted using an algorithm developed by King et al., 1978, to derive the size distribution of the boreal aerosols. Mie theory will then be used to calculate the aerosol phase function and single-scattering albedo.  Finally, the atmospheric correction algorithm of Wrigley et al., 1992, will be used to atmospherically correct selected NS001 Thematic Mapper (TMS), Landsat Thematic Mapper (TM), and Moderate-resolution Imaging Spectrometer (MODIS) Airborne Simulator (MAS) data collected during the 1994 BOREAS IFCs 1-3.

Atmospheric correction of Landsat TM and other satellite data will use the aerosol properties derived from surface optical depth measurements. Atmospheric correction of NS001 and MAS data will use aerosol properties derived from the airborne optical depth measurements as well as those from the surface measurements.

1.6 Related Data Sets

BOREAS RSS-12 Airborne Tracking Sunphotometer Measurements

BOREAS RSS-11 Ground Network of Sunphotometer Measurements

BOREAS RSS-18 Ground Sunphotometer Measurements in the SSA

2. Investigators

2.1 Investigators Names and Titles

Robert C. Wrigley (retired 1995)

Principal Investigator

Co-Investigators:

Michael A. Spanner

Robert E. Slye

Philip B. Russell

John M. Livingston

2.2 Title of Investigation

BOREAS Aerosol Determinations and Atmospheric Correction for BOREAS Imagery

2.3 Contact Information

Contact 1

--------------------------------

Michael A. Spanner

Co-Investigator

Johnson Controls World Services

NASA Ames Research Center

Moffett Field, CA

(415) 604-3620

(415) 604-4680 (fax)

mspanner@gaia.arc.nasa.gov

Contact 2

--------------------------------

Jaime Nickeson

NASA/GSFC

Greenbelt, MD

(301) 286-3373

(301) 286-0239 (fax)

Jaime.Nickeson@gsfc.nasa.gov

3. Theory of Measurements

The instrument measures direct beam solar radiation for 10 channels in the visible and near infrared wavelengths. The solar radiation data are collected in the form of voltages. The instrument was calibrated both before and after the experiment using the Langley plot technique.  For calibration, data are collected at a number of solar angles from low solar elevation (air mass = 5) to high solar elevation (air mass = 1.8). A regression is developed between log voltage and air mass. This regression equation is then extrapolated to an air mass of 0.  This value, called the zero air mass intercept voltage, is the value used to calibrate of the instrument in a given channel. Great care must be taken to ensure the stability of these intercept voltages over time.  A calibration history is maintained that attests to the stability of the instrument.  The voltages measured by the instrument during the BOREAS IFCs were converted to total optical depth using the zero air mass intercept voltages calculated from the calibrations using the equation:

V/V0 = (Rm/R)2 exp(-mt)

where V is the measured voltage, V0 is the zero air mass voltage intercept, R is the radius of Earth's orbit at the time, Rm is the mean radius, m is the air mass at the time, and t is the total optical depth (usually written as the Greek letter tau). The aerosol optical depth is calculated from the total optical depth by subtracting a number of components that contribute to the total optical depth: Rayleigh scattering and gaseous absorption due to ozone and NO2. The Rayleigh optical depth is calculated using pressure measured on the aircraft. NO2 and ozone optical depths are subtracted from the total minus Rayleigh optical depth to obtain the aerosol optical depth. NO2 abundance is obtained from climatology tables based on Noxon, 1979, and convolved with absorption coefficients at field sunphotometer wavelengths. Ozone optical depth is calculated using ozone abundances from the Total Ozone Mapping Spectrometer (TOMS) satellite instrument and convolved with absorption coefficients at the field sunphotometer wavelengths.  The result of this processing is the aerosol optical depth measured in nine channels (not including the 940 nm water vapor channel) at approximately 1-minute intervals on the ground.

The correction of remote sensing data acquired from satellites or aircraft for effects due to the intervening atmosphere has proven to be a difficult problem. Not only does the atmosphere reduce the transmission of the incoming, reflected, and emitted radiation, but it contributes reflected and emitted radiation of its own. Under high aerosol concentration conditions, atmospheric radiation comprises over 90% of the satellite observed radiance, but even much smaller effects would degrade the quantitative use of these data unless they are taken into account. The interaction of radiation with the atmosphere is complex and has proved difficult to calculate without reference to measurements made at, or close to, the time and location of interest. Effects due to Rayleigh scattering from atmospheric gases are well understood because the major gases (nitrogen, oxygen) that comprise 99% of the atmosphere are well mixed and their concentrations with altitude are known.  The effects due to small particle (aerosol) scattering are quite variable because of the wide range of aerosol concentrations and the variety of aerosols found in the atmosphere. Because aerosol concentrations cannot be known a priori, they must be measured at the time and location of remote sensing data acquisition.

The physical properties of aerosols such as size, shape, refractive index, and concentration in the atmosphere control the aerosol interaction with light according to a set of optical properties. Three fundamental properties are (1) the aerosol optical depth, an indirect measure of the size and number of particles present in a given column of air; (2) the single scattering albedo, the fraction of light intercepted and scattered by a single particle; and (3) the phase function, a measure of the light scattered by a particle as a function of angle with respect to the original direction of propagation.

4. Equipment

4.1 Sensor/Instrument Description

The automated solar radiometer instrument consists of a 10-channel solar radiometer, solar-tracking mount, and data acquisition/controller box with tracking and temperature control.

4.1.1 Collection Environment

The data collection took place at either the Sandy Bay Campground site (main site) or the parking lot at the Ship’s Lantern Hotel, both in Candle Lake, Saskatchewan.  Data were taken to coincide with NASA C-130 flights where airborne sunphotometer measurements were being taken.  Ground measurements were more frequent than aircraft flights however.  The field sunphotometer collected data on 52 days between 25-May-1994 and 19-Sep-1994.

4.1.2 Source/Platform

The instrument is mounted on a short tripod that rests on the ground.

4.1.3 Source/Platform Mission Objectives

The AGSP was developed to obtain accurate multispectral atmospheric extinction measurements in the field for the overall purpose of atmospheric correction of remotely sensed data.

4.1.4 Key Variables

The primary quantity being measured is the total optical depth.  The aerosol optical depth is derived by subtracting optical depths caused by other components of the atmosphere: Rayleigh scattering, ozone absorption, and NO2 absorption.

4.1.5 Principles of Operation

The instrument measures energy in the direct beam of the Sun. From the calibrations developed before and after the experiment, these voltages are converted to aerosol optical depth, which is a measure of the extinction of the direct solar beam by aerosols and particulates in the atmosphere.

4.1.6 Sensor/Instrument Measurement Geometry

The field sunphotometer has a 2.0 degree field of view (FOV) and is heated to 44 °C to maintain temperature stability.  It has 10 filters. The nominal wavelengths and the full width half maximum (FWHM) for the instrument are presented in the following table.

Wavelength (nm)  FWHM (nm)

---------------  ---------

380.2          11.7

401.1          10.2

438.6          10.6

521.6          11.6

608.5          10.2

666.9          10.6

779.3          10.1

865.6          12.6

939.8          11.6

1027.1           7.4

4.1.7 Manufacturer of Sensor/Instrument

Dr. John Reagan

Department of Electrical and Computer Engineering

University of Arizona

Tucson, AZ

(520)621-6203

4.2 Calibration

4.2.1 Specifications

Factors that may affect calibration are instrument variations that may occur between calibrations. Significant drifts in calibration during the time period of the experiment were not observed.

4.2.1.1 Tolerance

The aerosol optical depths are accurate to the uncertainties given with the data.

4.2.2 Frequency of Calibration

The instrument was calibrated at the Mt. Lemmon Steward Observatory, Tucson, AZ, in April 1994 (before the field season) and at the Mauna Loa Observatory, HO, in November 1994 (after the field season).

4.2.3 Other Calibration Information

5. Data Acquisition Methods

The instrument is initialized through the controller box and then a solar radiometer telescope is manually aligned with the solar image in the crosshairs. Once the telescope is aligned, the automated solar radiometer will collect data at the selected time interval by first tracking the Sun, then reading the output of all 10 channels, temperature, and the time of data collection. The time interval used for BOREAS was 1-minute. After each data collection sequence, the solar radiometer telescope is stepped away from direct solar alignment to reduce solar exposure on the interference filters. The data collection continues until the final stop time is reached or data collection is terminated. At this point, the instrument turns off the heating elements and is ready to transmit data to a computer through the RS232 port. The instrument has nonvolatile memory.

6. Observations

6.1 Data Notes

None given.

6.2 Field Notes

The field sunphotometer operator normally takes notes of significant events while the instrument is acquiring data. These notes supplement the data file of detector voltages or optical depths and permit determination of the presence of variable cloud interference with remote sensing data collection. The notes, if any, help identify data problems during processing. Anyone interested in these notes should contact RSS-12 personnel at NASA ARC.

7. Data Description

7.1 Spatial Characteristics

The field sunphotometer views the Sun with a 2-degree FOV and typically acquires data every minute during operation. The system is not moved during a collection period, which was about 4 hours at the main site, Sandy Bay. A collection period is a continuous data acquisition cycle.

7.1.1 Spatial Coverage

The field sunphotometer was operated from two locations in the Southern Study Area (SSA): Ship's Lantern Hotel and Sandy Bay Campground.

BORIS (X,Y)       UTM (E,N)        Latitude,Longitude

-------------  ------------------  --------------------

Ship’s Lantern Hotel  377.02,321.47  481967.2,5955747.4  53.75005 N, 105.27347 W

Sandy Bay Campground  372.84,326.53  478233.5,5961137.1  53.79835 N, 105.33047 W

The NAD83 corner coordinates of the SSA are:

Latitude     Longitude

--------     ---------

Northwest    54.321 N     106.228 W

Northeast    54.225 N     104.237 W

Southwest    53.515 N     106.321 W

Southeast    53.420 N     104.368 W

7.1.2 Spatial Coverage Map

Not available.

7.1.3 Spatial Resolution

The field sunphotometer views the Sun with a 2-degree FOV.

7.1.4 Projection

Not applicable.

7.1.5 Grid Description

Not applicable.

7.2 Temporal Characteristics

7.2.1 Temporal Coverage

The AGSP typically acquires data once every minute during operation. The system was not moved during a collection period.

Data were acquired during three IFCs in 1994. The data were intended to be coincident with the aircraft and satellite overpasses. The days, times, and locations were:

Date         Time (UTC)         Location

-----------  -----------------  ---------

25-May-1994  15:55:09-21:31:09  Sandy Bay

26-May-1994  15:13:17-19:00:13  Sandy Bay

27-May-1994  14:30:07-20:00:13  Sandy Bay

29-May-1994  14:27:08-18:50:08  Sandy Bay

31-May-1994  14:41:09-21:00:08  Sandy Bay

31-May-1994  11:42:11-13:13:08  Ship's Lantern

01-Jun-1994  14:19:10-21:00:09  Sandy Bay

04-Jun-1994  13:50:10-21:00:10  Sandy Bay

06-Jun-1994  15:17:08-19:00:10  Sandy Bay

06-Jun-1994  12:15:09-14:40:20  Ship's Lantern

07-Jun-1994  12:03:07-21:00:10  Ship's Lantern

08-Jun-1994  15:00:09-17:45:08  Sandy Bay

10-Jun-1994  12:02:13-16:35:08  Ship's Lantern

11-Jun-1994  15:51:08-19:00:19  Sandy Bay

11-Jun-1994  12:10:09-12:40:07  Ship's Lantern

20-Jul-1994  15:52:05-21:00:10  Sandy Bay

21-Jul-1994  15:35:11-22:30:08  Sandy Bay

21-Jul-1994  12:04:10-15:00:10  Ship's Lantern

22-Jul-1994  15:31:09-17:30:10  Ship's Lantern

23-Jul-1994  15:40:08-19:00:11  Sandy Bay

23-Jul-1994  11:58:09-14:40:09  Ship's Lantern

24-Jul-1994  15:30:08-21:30:10  Sandy Bay

24-Jul-1994  12:03:12-14:41:08  Ship's Lantern

25-Jul-1994  15:31:09-22:30:21  Sandy Bay

25-Jul-1994  12:12:11-14:39:12  Ship's Lantern

26-Jul-1994  15:43:10-19:00:10  Sandy Bay

27-Jul-1994  15:22:10-22:20:10  Sandy Bay

28-Jul-1994  16:09:09-19:00:09  Sandy Bay

30-Jul-1994  15:30:11-21:00:10  Sandy Bay

31-Jul-1994  15:36:11-19:00:14  Sandy Bay

01-Aug-1994  16:40:11-19:00:10  Sandy Bay

02-Aug-1994  15:47:12-20:22:12  Sandy Bay

04-Aug-1994  19:10:11-23:00:09  Sandy Bay

31-Aug-1994  18:38:05-22:00:09  Sandy Bay

01-Sep-1994  15:27:04-23:00:09  Sandy Bay

01-Sep-1994  13:04:09-14:45:09  Ship's Lantern

02-Sep-1994  15:30:04-23:00:09  Sandy Bay

02-Sep-1994  12:40:10-14:45:08  Ship's Lantern

05-Sep-1994  15:25:03-22:02:12  Sandy Bay

05-Sep-1994  12:45:07-14:44:08  Ship's Lantern

06-Sep-1994  15:22:03-23:00:09  Sandy Bay

06-Sep-1994  12:43:09-14:45:08  Ship's Lantern

07-Sep-1994  15:44:04-22:31:08  Sandy Bay

07-Sep-1994  12:54:09-14:30:08  Ship's Lantern

12-Sep-1994  17:06:05-20:06:03  Sandy Bay

13-Sep-1994  13:21:07-22:00:09  Sandy Bay

14-Sep-1994  18:35:06-19:26:09  Sandy Bay

15-Sep-1994  19:43:04-23:00:08  Sandy Bay

16-Sep-1994  13:35:04-23:00:19  Sandy Bay

17-Sep-1994  13:35:03-23:00:08  Sandy Bay

18-Sep-1994  13:31:05-23:00:18  Sandy Bay

19-Sep-1994  13:44:04-22:30:09  Sandy Bay

7.2.2 Temporal Coverage Map

Not available.

7.2.3 Temporal Resolution

The AGSP typically acquires data once every minute during operation. The data collection itself normally takes 10 or 15 seconds, dominated primarily by the time needed to re-point the instrument for solar tracking.

7.3 Data Characteristics

Data characteristics are defined in the companion data definition file (r12sunpd.def).

7.4 Sample Data Record

Sample data format shown in the companion data definition file (r12sunpd.def).

8. Data Organization

8.1 Data Granularity

All of the Automated Ground Sunphotometer Measurements in the SSA Data are contained in one dataset.

8.2 Data Format(s)

The data files contain numerical and character fields of varying length separated by commas. The character fields are enclosed with  a single apostrophe marks. There are no spaces between the fields.  Sample data records are shown in the companion data definition files (r12sunpd.def).

9. Data Manipulations

9.1 Formulae

For all sunphotometer channels except the 940 nm, the Bouguer-Lambert-Beer extinction law was used to describe the attenuation of solar radiation:

V = (R'/R)2 V0 exp(-m tau) = V'0 exp(-m tau)

where V is the output voltage of the detector at a given wavelength, V0 is the zero air-mass voltage intercept at that wavelength for the mean Earth-Sun separation R', R is the Earth-Sun separation at the time of observation, m is the atmospheric air mass between the instrument and the sun, tau is the wavelength-dependent total vertical optical depth above the sunphotometer, and V'0 is the zero-air-mass voltage intercept for the Earth-Sun separation R at the time of observation. The 940-nm channel requires different processing and is not included this data set.

The logarithm of the above equation,

ln V = ln V'0 - m tau,

is used in calibration to provide the V'0 values for each channel (i.e., zero air mass Langley plot intercept voltages). When the detector voltages are plotted against the air mass, the intercept is the V'0. After calibration, this equation can be solved for tau to provide the total optical depth. The total optical depth is then decomposed using

tau = tau_r + tau_a + tau_O3 + tau_NO2 + tau_H2O,

where these terms are the optical depth due to Rayleigh scattering, aerosols, ozone, NO2, and water vapor, respectively. The source for each of these terms is given in Section 7.3. Water vapor was ignored because it contributes only in the 940-nm channel.

This description is taken from Spanner et al., 1990, where more information concerning the data processing can be found.

9.1.1 Derivation Techniques and Algorithms

Description of algorithms can be found in Spanner et al., 1990.

9.2 Data Processing Sequence

9.2.1 Processing Steps

The steps for processing are as follows: 1) acquire the data; 2) transfer data to computer; 3) run program to reformat data; 4) run a program to calculate all the variables, including solar zenith angle, air mass, Rayleigh optical depth, and instantaneous optical depth (total optical depth minus Rayleigh optical depth); 5) calculate NO2 and ozone optical depths from Noxon et al., 1979, and TOMS data, respectively; and 6) subtract NO2 and ozone to derive aerosol optical depth.

The ozone abundance was determined from the TOMS satellite instrument convolved with ozone absorption coefficients from Penney (1979). The following table shows the values calculated for NO2 and ozone optical depth, which were subtracted from the instantaneous optical depth to derive the aerosol optical depth.

Wavelength  NO2 Tau  Ozone Tau

380       0.003     0.000

401       0.003     0.000

439       0.002     0.001

522       0.001     0.015

608       0.0       0.041

667       0.0       0.014

779       0.0       0.002

866       0.0       0.001

1027       0.0       0.0

9.2.2 Processing Changes

The processing sequence has not changed over time.

9.3 Calculations

9.3.1 Special Corrections/Adjustments

No special corrections or adjustments have been made.

9.3.2 Calculated Variables

Description of algorithms can be found in Spanner et al., 1990.

9.4 Graphs and Plots

Plots have been provided to BORIS and can be made available upon request.

10. Errors

10.1 Sources of Error

Calibration errors are the main source of error in the derivation of aerosol optical depth.

10.2 Quality Assessment

10.2.1 Data Validation by Source

Data were compared with the RSS-11 ATSP measurements (see related data sets, Section 1.6).

10.2.2 Confidence Level/Accuracy Judgment

The data are of high quality, because a good calibration of the instrument was performed before and after the BOREAS field collection effort. However, post-BOREAS calibration data were not available for the 1027-nm channel.

10.2.3 Measurement Error for Parameters

Uncertainties for the aerosol optical depths were determined by using uncertainty propagation through the algorithm.  The aerosol optical depth uncertainty is dependent on the uncertainty in the Rayleigh, ozone, and NO2 optical depths, as well as the uncertainty in the intercept voltage (calibration error), instantaneous measurement, and airmass.  Aerosol optical depth uncertainties are given in the data files and are summarized in Section 7.3 of this document.

10.2.4 Additional Quality Assessments

None.

10.2.5 Data Verification by Data Center

Visual review and use of selected subsets of the data have shown them to be of good quality with no noteworthy problems.

11. Notes

11.1 Limitations of the Data

11.2 Known Problems with the Data

Because the post-BOREAS calibration in November 1994 did not provide intercept voltage data for the 1027-nm channel, an updated voltage was not available. An estimate based on previous calibrations was used.

11.3 Usage Guidance

The values of aerosol optical depth are accurate instantaneous values of aerosol optical depth. These data were taken every minute; therefore, under conditions of rapid variability in cloudiness or haze, the data may not be internally consistent or appropriate. It is useful to calculate averages of aerosol optical depth over periods of time (for example, 30-minutes) to get a more accurate measure of the average conditions at a site.

11.4 Other Relevant Information

The aerosol optical depth at 940 nm was not calculated because this channel primarily measures absorption due to water vapor.

12. Application of the Data Set

These data can be used for correcting various visible and infrared satellite and aircraft image products or for characterizing the atmospheric aerosols at the times of the flights.

13. Future Modifications and Plans

None.

14. Software

14.1 Software Description

NASA ARC software was developed in FORTRAN on a VAX to implement the data processing procedure described in Section 9.1. Input data include sunphotometer data files as well as ozone and NO2 optical depth parameters. Aerosol optical depths were calculated and written to the data files.  No special software is needed to read the data files because they are stored comma delimited.

14.2 Software Access

This software is used to generate the data product from the detector voltages and is not needed to use the data.

15. Data Access

15.1 Contact Information

Ms. Beth Nelson

BOREAS Data Manager

NASA GSFC

Greenbelt, MD

(301) 286-4005

(301) 286-0239 (fax)

beth@ltpmail.gsfc.nasa.gov

15.2 Data Center Identification

See Section 15.1

15.3 Procedures for Obtaining Data

Users may place requests by telephone, electronic mail, or fax.

15.4 Data Center Status/Plans

The RSS-12 ground sunphotometer data are available from the Earth Observing System Data and Information System (EOSDIS), Oak Ridge National Laboratory (ORNL), Distributed Active Archive Center (DAAC). The BOREAS contact at ORNL is:

ORNL DAAC User Services

Oak Ridge National Laboratory

Oak Ridge, TN

(423) 241-3952

ornldaac@ornl.gov

ornl@eos.nasa.gov

16. Output Products and Availability

16.1 Tape Products

None.

16.2 Film Products

None.

16.3 Other Products

The data are available as tabular ASCII files.

17. References

17.1 Platform/Sensor/Instrument/Data Processing Documentation

Portable Radiometer Data Reduction Manual for use with PDATA7.

17.2 Journal Articles and Study Reports

Bruegge, C.J., R.N. Halthore, B. Markham, M. Spanner, and R. Wrigley. 1992. Aerosol optical depth retrievals over the Konza prairie. Journal of Geophysical Research 97(D17):18743-18758.

King, M., D. Bryne, B. Herman, and J. Reagan. 1978. Aerosol size distributions obtained by inversion of spectral optical depth measurements. J. Atmos. Sci. 35:2153-2167.

Noxon, J. 1979. Stratospheric NO2, 2, Global behavior, J. Geophys. Res. 84:5067-5076.

Penney, C.M. 1979. Study of temperature dependence of the Chappuis band absorption of ozone, NASA Contract Rep. 158977, General Electric  Company, Schenectady, N.Y.

Russell, P., J. Livingston, E. Dutton, R. Pueschel, J. Reagan, T. DeFoor, M. Box, D. Allen, P. Pilewskie, B. Herman, S. Kinne and D. Hofmann. 1994. Pinatubo and pre-Pinatubo optical depth spectra:  Mauna Loa measurements, comparisons, inferred particle size distributions, radiative effects, and relationship to lidar data. J. Geophys. Res. 98:22,969-22,985.

Sellers, P., and F. Hall. 1994. Boreal Ecosystem-Atmosphere Study: Experiment Plan. Version 1994-3.0, NASA BOREAS Report (EXPLAN 94).

Sellers, P., and F. Hall. 1996. Boreal Ecosystem-Atmosphere Study: Experiment Plan. Version 1996-2.0, NASA BOREAS Report (EXPLAN 96).

Sellers, P., F. Hall, and K.F. Huemmrich. 1996. Boreal Ecosystem-Atmosphere Study: 1994 Operations. NASA BOREAS Report (OPS DOC 94).

Sellers, P., F. Hall, and K.F. Huemmrich. 1997. Boreal Ecosystem-Atmosphere Study: 1996 Operations. NASA BOREAS Report (OPS DOC 96).

Sellers, P., F. Hall, H. Margolis, B. Kelly, D. Baldocchi, G. den Hartog, J. Cihlar, M.G. Ryan, B. Goodison, P. Crill, K.J.Ranson, D. Lettenmaier, and D.E. Wickland. 1995. The boreal ecosystem-atmosphere study (BOREAS): an overview and early results from the 1994 field year. Bulletin of the American Meteorological Society. 76(9):1549-1577.

Sellers, P., and F. Hall. 1997. BOREAS Overview Paper. JGR Special Issue (in press).

Spanner, M., R. Wrigley, R. Pueschel, J. Livingston, and D. Colburn. 1990. Determination of atmospheric optical properties for the First ISLSCP Field Experiment (FIFE). Journal of Spacecraft and Rockets 27:373-379.

Wrigley, R.C., M.A. Spanner, R.E. Slye, R.F. Pueschel, and H.R. Aggarwal. 1992. Atmospheric correction of remotely sensed image data by a simplified model. Journal of Geophysical Research 97(D17):18797-18814.

Young, A. 1980. Revised depolarization corrections for atmospheric extinction. Applied Optics 19:3427-3428.

17.3 Archive/DBMS Usage Documentation

18. Glossary of Terms

air mass       - secant of the solar zenith angle

optical depth  - an indirect measure of the size and number of

particles present in a given column of air, which is a

measure of the extinction of the direct solar beam by

aerosols and particulates in the atmosphere, or by

scattering.  Also referred to as optical thickness.

phase function - a measure of the light scattered by a particle as a

function of angle with respect to the original

direction of propagation

radiometer     - an instrument for measuring radiant energy

Rayleigh       - wavelength-dependent scattering directly proportional

scattering       to (1 + cos2 angle) and indirectly proportional

to wavelength4

single         - the fraction of light intercepted and scattered by a

scattering       single particle

albedo

19. List of Acronyms

AGSP    - Automated Ground Sunphotometer

ARC     - Ames Research Center

ATSP    - Airborne Tracking Sunphotometer

BOREAS  - BOReal Ecosystem-Atmosphere Study

BORIS   - BOReas Information System

DAAC    - Distributed Active Archive Center

EOS     - Earth Observing System

EOSDIS  - EOS Data and Information System

FIFE    - First ISLSCP Field Experiment

FOV     - Field of View

FWHM    - Full Width Half Maximum

GMT     - Greenwich Mean Time

GSFC    - Goddard Space Flight Center

IFC     - Intensive Field Campaign

ISLSCP  - International Satellite Land Surface Climatology

MAS     - MODIS Airborne Simulator

MODIS   - MODerate-resolution Imaging Spectrometer

NASA    - National Aeronautics and Space Administration

NSA     - Northern Study Area

ORNL    - Oak Ridge National Laboratory

PANP    - Prince Albert National Park

RSS-12  - Remote Sensing Science Team 12

SSA     - Southern Study Area

TM      - Thematic Mapper

TMS     - Thematic Mapper Simulator

TOMS    - Total Ozone Mapping Spectrometer

URL     - Uniform Resource Locator

UTC     - Universal Time Code

20. Document Information

20.1 Document Revision Dates

Written:         07-Jan-1997

Last Updated:    04-May-1998

20.2 Document Review Dates

BORIS Review:    19-May-1997

Science Review:  27-Jun-1997

20.3 Document ID

20.4 Citation

Please acknowledge the NASA ARC investigation (RSS-12) and Robert Wrigley, Principal Investigator, if these data are used or referenced. If appropriate, the references cited in Section 17 may be used.

20.5 Document Curator

20.6 Document URL

Keywords

Aerosols

Optical Depth

Ozone